Hitachi V-212 Oscilloscope # 6060378 Tests			2023-06-20
Performance tests:
Trace Rotation required as location changed from last setup
Channel1: All 10 ranges checked as well as x5 and variable gain. Operating correctly.
Channel2: All 10 ranges checked as well as x5 and variable gain. Operating correctly.
Timebase: All 19 ranges checked as functional as well as the Sweep Variable.
CH1, CH2 Mode: All settings checked, some minor DC unbalance on some ranges.
Timebase Mode: All options checked.
Auto Mode: Triggers on 0.3Div minimun as opposed to 0.5Div called for in specification page.
Note the beam intensity diminishes at less than 0.5Div.
Triggering tested from <10Hz to >2MHz.
External Trigger: Triggers with 100mV P-P.
Tested from <10Hz to >2MHz.
Note as with the Auto mode the beam intensity diminishes at less than 0.5Div.
Intensity: Brillance is not room brightening (as it shouldn't be) but normal for test bench operation. 
Both probes tested as having 10:1 attenuation. One had not been compensated to this scope.
No issue of concern arose during the extensive testing.
Noted the CH1,CH2 Mode switch knob is split but still fully functional.

Addendum  2023-06-21
Returned with protective plastic bag around scope missinghchh.
No wallets returned for probes nor data probe sheets.